# WOMEN’S HEALTHCARE ASSOCIATES, LLC

TITLE:	Decontamination and Sterilization of Instruments	 POLICY C-12

APPLIES TO:	Employees

POLICY:	Women’s Healthcare Associates, LLC (WHA) staff will follow the decontamination and sterilization process to ensure that instruments are thoroughly cleaned, packaged, and stored; that quality assurance is supported; and that the autoclave is properly maintained.

# PROCEDURE:	All clinical services staff members who participate in the decontamination and steam sterilization process within WHA are accountable for knowing the following:

- The basic process steps- Instrument Cleaning; Packaging and Preparation; Sterilization; Quality Assurance Monitoring; Instrument Storage; and Autoclave Maintenance.

- Required equipment list-

  - PPE, as detailed below

  - Enzyme detergent

  - Instrument cleaning brushes and sponges

  - Instrument lubricant and/or cleaner

  - Packaging material

  - Steam indicator tape

  - Steam indicator strips

  - Biological indicators

  - Thermometer for solution temperature

  - Spore log (either WHA or manufacturer log is acceptable)

  - Distilled water

  - Autoclave

  - WHA log for sterilizer loads

  - WHA temperature log for solution preparation

  - WHA log for instruments needed prior to quarantine release

Step One: Instrument Cleaning

- Required PPE for instrument cleaning includes disposable gloves covered with utility gloves, protective eyewear (goggles, face shield, or eyeglasses with side shields), head cover, mask, and waterproof gown. Sink splash guards are not a substitute for eye protection or head cover.  Use beard covers if indicated.

- Handle instruments carefully. Instruments should never be piled in a heap; this may cause damage to occur if instruments become entangled. Keep sharps separate from the rest of the instruments to avoid damaging or dulling the cutting edge or injuring the employee. Specialty instruments like scopes should be cleaned according to the manufacturer’s instructions, which are available on SharePoint

- Disassemble all instruments with multiple parts; hinged or jointed items should be in the open position during cleaning and disinfecting.

- Rinse instruments with tap water to remove gross debris such as blood, body fluids, and tissue.

- After rinsing, the instruments should be dried before being placed in the solution of enzymatic detergent. Leave instruments on a towel to air dry or manually dry with a paper towel or lint free towel. Lint-free towels must be used when required by the manufacturer’s instructions. Drying the instruments prevents the dilution of the disinfectant when they are placed in the solution.

- Prepare an enzymatic detergent bath, ensuring that the solution reaches the starting preparation temperature required by the manufacturer’s directions. Soak instruments in the solution of enzymatic detergent for at least 2 minutes to remove any remaining debris. Never store instruments in the solution overnight.

- Lightly scrub instruments with a Teflon or nylon brush. It is also acceptable to use an instrument sponge (not a kitchen sponge) to aid in cleaning blunt instruments.

        - Brush care: Replace the brush once a month, or as needed for bristle damage. Decontaminate daily by placing in solution for 2 minutes, thoroughly rinsing, air drying, and then inspecting for bristle damage.

        - Sponge care: Discard and replace instrument sponges daily.

- Rinse instruments under warm water for at least 20-30 seconds to fully remove all of the detergent.

- Leave instruments on a towel to air dry or manually dry with a paper towel or lint free towel.

Step Two: Packaging and Preparation

- After cleaning, inspect the instruments closely:

  - Scissors- check for nicks, smooth gliding, and sharpness. Cut a piece of tissue (Kleenex), the tips should cut smoothly and release the tissue without catching. Check for loose screws.

  - Forceps- make sure the tips meet properly.

  - Hemostats- check that the tips meet and that serrations interlock when ratchets are closed. Check that jaws are aligned, and that ratchets will lock and hold in all three positions.

  - Needle holders- jaws should show no light between the serrations when held up to a light source.

  - All instruments with box locks- check for cracks in the locks.

  - Instrument tape is acceptable to use if it is inspected after each cleaning and remains in good repair.

  - If repairs need to be considered, follow the current workflow  for equipment transferring and tracking. Any instrument that appears to need repair is to be autoclaved before being sent for repair.

- Protect the tips of sharp instruments with a tip protector or sleeve protector approved for autoclave use.

- Make sure all hinged instruments are in an open position; sleeve protectors, instrument guards, and some tip protectors will help maintain the open position. Steam will not reach all areas of instruments autoclaved in a closed position.

- Use clinic-approved wraps, pouches and containers to package items for sterilization.

- Packs:

  - Start with two blue sterile wraps. Lay the instrument tray on the wraps and then add the instruments, gauze, etc. It is not necessary to place layers of gauze between instruments.

  - Insert a steam indicator strip inside the center of each pack with the instruments. This indicates if the steam from the autoclave has fully penetrated the inside of the pack.

  - To enclose the pack, wrap with the first layer followed by the second. Ensure that the opening is folded appropriately for easy retrieval of the instruments. Apply steam indicator tape to seal the pack and write the item name, full date, expiration date, batch number, autoclave location (when indicated), and your initials.

  - When sterilizing metal containers, do not place a lid on the container. Steam will not be able to penetrate to ensure sterilization. Place the lid under the container and wrap the lid and container with separate blue sterile wraps.

  - Small packs are best because steam penetrates them more easily. A pack should be no larger than 12 X 12 X 20 inches.

- Pouches:

  - Place instruments in pouches with the handles down, so that when the package is opened the handle will be exposed.

  - Never punch, or poke instruments through the wrapping to open the pouch. This can damage the instrument.

  - Add a steam indicator strip to the pouch.

  - Write the full date, expiration date, batch number, location of autoclave (when indicated), and your initials on the pouch.

- Unwrapped Instruments: Add steam indicator strip on the tray including the full date, batch number, location of autoclave (when indicated), and your initials.

- Other considerations: Wrap glass and delicate instruments individually. Position rubber items to prevent contact with metal.

Loading the autoclave:

All items must be loaded into the autoclave chamber in a position that allows for free circulation and penetration of steam. Best practices to reduce pooling of moisture and improve drying efficiency include:

- Process unwrapped instruments, packs, and pouches in separate loads whenever possible.

- Place items loosely inside the chamber; maintain approximately 1 to 3 inches between small packs and 2 to 4 inches between large packs. Leave at least 1 inch of space between packs and nearby autoclave trays.

- Packs containing layers of fabric, such as dressings, should be placed in a vertical position.

- Pouches should be positioned on their sides, in matching orientation (i.e. the plastic sides of each pouch should be facing the same direction). If not positioned on their sides, the pouches may be placed in a single layer with plastic side up.

- Unwrapped instruments must be placed in a single layer on autoclave trays.

- Jars and glassware must be placed on their sides with their lids removed.

Step Three: Sterilization

- Prior to autoclaving the instruments, LOG them on the sterilizer log sheet. This information must include date and time, brief descriptor/number of packs, batch number, and your initials.

        - The person taking out the load will need to initial the log sheet for the batch they remove and record the time of unloading.

- For detailed operational instructions, refer to the manufacturer's instructions for your specific autoclave.

- The Midmark Steam Sterilizers at WHA use the programmed cycles outlined in Table 1.

- Refer to Policy C-39 for sterilization of pessaries and diaphragms.

- Follow manufacturer’s instructions for cryo tip cleaning and sterilization.

Table 1:

| Cycle Type | Cycle Button | Cycle Function |
| --- | --- | --- |
| Unwrapped |  | Process instruments loose on a tray, open glass or metal canisters, or other loose items at: 270° F (132° C) for 3 min / 30 min drying cycle. |
| Pouches |  | Process instruments in pouches or loosely wrapped at: 270° F (132° C) for 5 min / 30 min drying cycle. |
| Packs |  | Process packs of instruments or textiles at: 250° F (121° C) for 30 min / 30 min drying cycle. |
| Handpieces |  | Process dental handpieces. (Not used at WHA) 270° F (132° C) for 6 min / 30 min drying cycle. |

Step Four: Quality Assurance Monitoring

The threat of infections caused by inadequate sterilization makes the monitoring process extremely important. There are three separate indicators to monitor the entire process of sterilization:

- Exposure control- Steam indicator tape

- Pack control- Internal steam indicator strips

- Load control- Biological indicators (spore testing)

- Close all packs with steam indicator tape. Diagonal marks on the tape change color when the outside of the pack has been exposed to steam sterilization of the correct duration and temperature.

- Insert a steam indicator strip into the middle of each sterile pack. The “accept” part of the indicator darkens when the inside contents of the pack have been exposed to steam sterilization.

  - Internal steam indicator strips should be placed in the center of the autoclave with every load of instruments.

  - If the strip does not change color to show adequate sterilization, run the load again with a new strip. If the second load fails, contact your on-site Clinical Services supervisor.

- Use biological indicators (spore testing) at least daily when an autoclave is used that day. Rapid biological indicators should be used with each load.

  - If instruments are not autoclaved daily, spore testing will occur weekly, or whenever an instrument, pack, or pouch is next autoclaved.

  - 3M biological indicators can be tested in a load that also has instrument packs.

  - Place the biological indicator in the most challenging area for the sterilizer- this is near the door on the bottom shelf for Midmark Steam Sterilizers.

  - Log the results of the Biological Indicator Strip on the WHA log sheet or Manufacturer log.

- Quarantine of instruments awaiting biological indicator results:

  - Instruments processed in the autoclave must be quarantined until the day’s spore test returns as passing.

  - When quarantine is not possible, instruments used prior to release from quarantine will be placed on a tracking log alongside the patients for whom the instruments are utilized. Patients found to be at risk for pathogen exposure will be offered appropriate follow-up.

- Response to positive biological indicator results:

  - If the biological indicator test fails, any products sterilized during the failed event period must be recalled and re-sterilized.

  - Document this action in the logbook.

  - Try to find the potential reason for sterilization failure prior to re-running the load. Refer to the Troubleshooting Chart in manufacturer’s instructions for use.

  - Re-run the load with another biological indicator. Do not release instruments until the second spore test has resulted and does not show repeat failure. It is acceptable to reprocess instruments in another autoclave. If two consecutive spore tests have failed, contact your supervisor for further evaluation.

Step Five: Instrument Storage

Instrument storage areas should be a minimum of 8-10 inches off the floor, 18 inches from the ceiling, and 2 inches from internal walls where possible. Closed drawers or cupboards are preferred.

- Do not stack packs or pouches. They should be stored in a single layer or in a vertical position.

- The area must remain clean and dry. If packs are not turned over often (i.e. weekly) and are not stored in a closed drawer or cupboard, protect packs from microorganisms and moisture by covering them with plastic or placing them in a “zip-lock” bag.

- Use the first-in, first-out method when obtaining instruments from storage.

- Shelf Life: Items packaged and sterilized have a shelf life of three years past the sterilization date, unless an event occurs that breaks the sterile barrier or if the sterilization pouch has an expiration date listed. Such events include tears, punctures, moisture intrusion, or opening of the package.

Step Six: Autoclave Maintenance

Refer to the manufacturer’s instructions for use for illustrations and detailed guidance related to autoclave maintenance.

Daily:

- Wipe down the exterior of the sterilizer each day. Use only quaternary disinfectants such as Sani-Cloth AF3 (“gray top” wipes) to disinfect the unit. Other products can damage the unit. Follow the instructions provided with the cleaning product used.

- Clean the sterilizer door gasket with a damp cloth. Examine for possible damage.

Weekly:

- Disconnect the upper portion of the reservoir drain tube from the panel clips, bend it downward, and drain the water from the reservoir into a suitable container, e.g. a bucket, and dispose of the water.

- Remove trays, tray rack, and tray plate form the sterilizer.

- Wash trays, racks, plate and the inside of the sterilizer chamber mild soap or Speed-Clean and distilled water.

- Refill reservoir with distilled water.

Monthly:

Clean Chamber / Plumbing

- With a cooled chamber, drain the reservoir and refill with clean distilled water. Add one ounce of Speed-Clean sterilizer cleaner directly to the bottom of the chamber.

- Run one Pouches cycle. Press the Stop button when Dry Cycle begins. (Dry Cycle is not needed during maintenance.)

- Drain reservoir and refill a second time with clean distilled water.

- Rinse by running one Unwrapped cycle. Push the “Stop” button when the drying cycle begins.

- Drain and refill reservoir with clean distilled water, then allow sterilizer to cool.

- Remove trays and tray rack. Wipe with a damp cloth.

- Remove and clean filters. The filters are intended to prevent debris from causing valve failures. Between regular monthly cleanings, if the fill or vent times become too long or items will not dry, the filters should be cleaned.

- Grasp the filter and gently pull away from chamber wall while twisting slightly.

- Clean filters with Speed-Clean and distilled water and a small stiff bristled brush. Rinse filters with distilled water. Replace filter(s) if debris cannot be removed by cleaning.

- Wipe out the inside of the chamber.

- Install filters. (Press inward, toward chamber wall while twisting slightly).

- Install tray plate, rack, and trays.

Remove / Clean Door & Dam Gaskets

- Remove door and dam gaskets from chamber door and then remove the gasket ring from the door gasket.

- Clean gaskets and rinse with Speed-Clean, distilled water, and a soft brush.

- Inspect gaskets for damage / shrinking / swelling. Replace gaskets if damage is apparent.

- Press the gasket ring into the channel in the door gasket and reinstall the gasket in the door. Reinstall the dam gasket.

Check Pressure Relief Valve

- Caution: During the pressure relief valve check, steam will be vented from under the sterilizer. To keep from being burned, place a steam barrier using rolled up towels around the bottom of the sterilizer.

- Select Unwrapped button.

- Press Start button.

- Wait until the pressure reaches 20 PSI (138 kPa).

- Next, pull upward firmly on the pressure relief lever for approximately 3 seconds, then release.

  - Steam should discharge freely from beneath the rear of unit when lever is pulled.  If the valve does not close completely when lever is released, pull the lever again and release quickly so that it snaps closed.

  - Repeat this until valve seats properly

- Press the Stop button. This aborts the cycle to prevent overheating.

Annually:

- The autoclave is to be checked by licensed Biomedical personnel to ensure that the autoclave is calibrated for temperature and pressure.

- Documentation of annual checks is kept on file with equipment maintenance records.

| Chief Executive Officer C-12 |  | Chief Medical Officer C-12 |
| --- | --- | --- |

| Chief Operations Officer C-12 |  |
| --- | --- |

FORMULATED	May 30, 2007

REVISED		February 3, 2010

REVISED 	June 25, 2012

REVISED		December 3, 2012

REVISED		January 3, 2013

REVISED		April 21, 2016

REVIEWED	August 20, 2018

REVISED		August 30, 2021

REVISED		April 10, 2023

REVISED		December 12, 2023

REVISED		August 7, 2025

REVISED		November 6, 2025